## Sheet1
MALDI TOF-TOF Tissue Imaging System under buy back of Existing bruker maldi ultraflex | Criteria Matching (Yes/NO) | Remarks
The MALDI Time of Flight Mass Spectrometer with highest quality performance in Mass resolution, Sensitivity and Mass accuracy for the below Applications and the workflows.
a) Tissue/MALDI Imaging of Metabolites, Drug Distribution, Proteins& Peptides and Lipids.
b) Identification of intact protein, protein/peptide identification and characterization, protein biomarker research, PTM identification and analysis, de novo sequencing, gel based and non-gel based applications, biopolymer analysis.
c) System should be highly suitable for the Analysis of Oligonucleotides, Nucleic acids, PNA-DNA Hybrids.
Technical Specifications:
i. Gridless MALDI source for maximum ion transmission and sensitivity.
ii. Scanning laser beam with 2 kHz repetition rate and modulated laser beam profile for best sensitivity.
iii. Laser beam diameters of 5 µm diameter for versatile applications with different sample preparation techniques; including variable power attenuator
iv. Different Laser focus patterns adapt for optimal sensitivity using different image resolution requirements for full pixel coverage from small to larger pixel size
v. True pixels at precise quadratic shape, adaptable to different resolution requirements down to 50 µm @ 20 Hz acquisition speed of a restricted area on a microscopic slide.
vi. Solid state laser with laser pulse energy ≥ 100 uJ.
vii. Guaranteed number of laser shots ≥ 3.5m
viii. 24/7 operation: Source cleaning is simple. Convenient ion source cleaning procedure without the use of any tools:
ix. Pulsed ion extraction technology for high mass accuracy and unmatched resolution spectra across an extended mass range
Performance Specifications
i. Linear mass range at least up to 300 kDa.
ii. High mass resolution over a wide mass range. In reflectron mode, resolution of 15,000 – 26,000 FWHM simultaneously across a wide m/z range 900 – 4500 in a seamless mass spectrum.
iii. MS resolution › 26,000 FWHM achieved in MS mode for m/z 3.147 and ≥ 15,000 FWHM for Cytochrome C (m/z 12.361)
iv. Reflectron mass accuracy ≤ 2.0 ppm with internal calibration
v. MS/MS sensitivity ≤ 250 amol [Glu1]-Fibrinopeptide B fragment 1,056 Da at S/N ≥100:1
Mass analyzer
i. True TOF/TOF technology with two stage acceleration and precursor ion selector betweenTOF1 and TOF2.
ii. Gridless linear and reflectron ion optics for detection in linear and reflectron mode, respectively.
iii. Ion optics adapts to analysis mode for best resolution and sensitivity: the MS/MS ion optics moves out the ion path for best MS sensitivity, the reflector is automatically adapted to MS vs MS/MS experiments for best resolution.
iv. CID (Collision-Induced Dissociation) Capability (e.g., Leu/Ile differentiation)
v. In source decay (ISD) for the direct fragmentation of intact proteins without the need for enzymatic digestion.
Free calibration for five years after the purchase
Data System and Software, WorkStation should have following technical features.
A) Molecular Tissue Imaging Applications.
Should Include the following:
An automatic sprayer with spraying booth, XY stage, spray nozzle, TMSP Control Software 4.1, Windows PC, switching valve, 5ml loop and necessary fittings, Isocratic Pump sprayer (10 ml pump head, stainless steel), installation and familiarization, with NE-300 trypsin digestion kit that includes NE-300 Pump, infusion only syringe pump. 1ml glass syringe, tubing and connections.
Dedicated Tissue image analysis software for mining large MALDI imaging data, statistical analysis and correlation to histological scans:
i. Supports the direct overlay of scanned histology microphotographs with the MALDI images to correlate histological features either with dedicated molecular signals or tissue state classifiers (disease state, etc.)
ii. User interface and performance to handle multiple datasets, e.g. for larger cohort analysis
iii. Univariate analysis: Mean intensities, standard deviations, uncovering discriminative m/z-markers (ROC analysis), hypothesis tests and discovery of – co-localized m/z values
iv. Unsupervised multivariate analysis: Spatial segmentation of large datasets, component analysis for extraction of underlying trends
v. Supervised tissue typing: Calculation of classifiers based on training data and classification of new observations / samples.
High performance computer (workstation) with adequate memory and all interfacing software and hardware. Computer specification should be submitted along with the technical bid.
a) The operating system should be based on Windows 10(or higher) (64bits) OS with adequate features to transfer data between computers for standalone data analysis
b) Suitable Laser printer should be provided
Other items
a) Standard two year warranty to be provided and additional 3 years warranty price to be quoted separately.
b) Essential target plates, Matrix & Standard for mass calibration should be provided
C) In case of any breakdown and will not be rectified within 10 days from the date of complaint log, warranty will be extend accordingly.
The system should be upgradable later for following:
Proteomics Applications
Software related to Protein identification and characterization, de novo sequencing, PTM’s discovery, Polymer analysis and LC-MALDI work flow must be available.
TLC – MALDI for LIPID analysis.
Software and online and offline database related accessories to perform the TLC MALDI Experiments should be available. Free upgrade to softwares and database for five years after the purchase
Training
15 days on-site training for two technicians from ICGEB and support to technicians for next 3 months