## Sheet1
Specifications of NIR Fluorescence Scanner | Creteria Matching (Yes/NO) | Remarks
Advanced and fully automated multipurpose NIR Fluorescence Scanner system to be able to document, perform absolute quantification and analyze the information from the following research applications:
· Infrared western blot, Multiplex Fluorescence Western blot,Fluorescence imaging, Colorimetry, Coomassive-stained Gels, RNA i analysis, Zymography, Cellular based Western assays, Protein Arrays, Non-isotopic EMSAs assay, Tissue section imaging, Whole organ imaging, DNA syto 60 gels, Multiplexing Imaging and ELISAs. All these application / usage by the system should be supported with Internationally peer viewed publications.
· Supplier ( Manufacturer /authorized distributor) should provide comprehensive imaging solutions which should include : Instrument, Reagents, Acquistion software, application/protocol support, hardware service support, Accurate data analysis software inline with current international journal guidelines.
System Hardware Specifications :-
· System which offers superior resolution of min. 21 um (micron ) or better and resolution claimed has to supported with proper documentation.
· Signal uniformity : Signal CV < 3 % across field of view
· System should have superior Multi- Laser based Optical Detection system that eliminates the need for post image capture data correction
· Wide Linear Dynamic range must be min. or greater than 6 orders/logs or 22 bit of magnitude for efficient detection and quantification of both faint and intense band on the same gel/blot as well as for all the chennels thereby offering optimal image acquisition in a single capture.
· Excitation System : System should use higher sensitive and superior solid state lasers diodes emitting at detection range of 680 to 690 nm and 780 to 790 nm respectively, must be the means of excitation of the fluorescent probes / labels/ dyes for high sensitivity & re-producible datas.
· Detection System : Detection of the fluorescent dyes must be in the ranges (710 nm - 730 nm) to (810 nm - 830 nm) for the 700 nm and 800 nm channels, respectively,
· No Excitation /detection thresholds below 680 nm – 700 nm will be accepted, in order to obtain a high signal-to-noise ratio. Detection system should preferably use avalanche Photodiodes (APD) for best Quantum Efficiency (QE at 800 nm: 8% for PMT, 80% for APD).
· Imaging Field size area should be minimum 25 cm x 25 cm or better
· Scanning speed should be more than 5 – 40 cms/second or better and focusing range adjustable 0 mm – 5 mm above the scan bed to obtain best signal to noise ratio.
· The system must utilize a sealed, flat-bed scanning surface to accommodate varying sample media (membranes, multi-well plates, gels, slides, etc.).
· Detection system should preferably based on Avalanche Photodiodes (APD) in contrast to photomultiplier tubes (PMTs) for best Quantum Efficiency (QE at 800 nm: 8% for PMT, 80% for APD).
· System should have the facility to include the calibration and verification plates optimized to confirm the consistent performance of the system for regulatory / publication purpose.
Image Acquisition & Data Analysis Software:
· User friendly software for image acquisition and precise analysis capabilities.
one-touch image acquisition everytime without a need for multiple exposures and also eliminates the need for post image capture data correction using image stacking/ flat fielding/ binning , providing uniformity across the entire area of imaging.
· Complete flexibility with automatic and manual detection of lanes and bands, using several algorithms.
· Data Analysis function should include 1D & 2D Quantitative analysis to find and labels lanes & bands, Molecular Weight calculation, Rf value calibration, Band intensities and Area Density calculation.
Acquisition Software: should have the following publication and grant submission benefits
Preserves unaltered raw data for publication or grant submission
Documentation of all image adjustments for publication or grant submission
Features Lab Notebook for accurate experimental record keeping
· Image data must be available for viewing in a minimum of three modes - colored overlay, colored single channel, and single channel grayscale. These three methods will help for a maximum amount of flexibility and detailed data analysis.
Data Analysis Software: should include following features for generating the highest quality Western blots data
Systematic workflow for user-to-user consistency
Validations and analysis workflows based upon recent publisher recommendations
Antibody validation
Combined Linear Range validation
Housekeeping Protein validation
Normalization
Replicate Analysis
Publication Package export
· Data and result output must be compatible with common spreadsheet and text software (e.g., Microsoft® Excel, Microsoft Word).
· System should be supplied with 10 User licence for Data analysis software for analysis purpose ( Not free downloadable software )
· Suitable Work Station with latest licenced windows 10 software along UPS should be given along with system
Sales after service Requirements :
· System – Specific Model should have been installed & used by atleast 10 leading research institutions or more in india for above mension applications with good sales after service support with maxium number of international publicaitons. List of installations of specific quoted model should be included in the quote.
· Quoted Model / specific instrument should have been “Proven” and be in the usage in leading research institutions ( CSIR /ICMR / DBT labs ) atleast for past 5 years or more and spares should be available for next 10 years of the instrument life.
· Min.5 Customer satisfactory / performance certificate for specific quoted model from the end user should be included in the quote.
· The manufacturer or their authorized distributor should provide highly knowledgeable, professional, and readily-accessible Field Application Scientists/ Factory trained Service engineers for on-site and / or remote training, servicing, and hardware, software, and application support.
· The manufacturer or their authorized distributor should support and supply customers with up-to date protocols and application guides for various applications available with the quoted models.
· If any specific consumables is required, they should be supplied along with the machine equivalent to first one year uses. Also, the rates for such consumables should be fixed for next 3 years.
· Warranty- 5 years with all spares