## Sheet1
Specifications for Vertical Autoclave | Creteria Matching (Yes/No) | Remarks
1. Chamber volume should be 90-100 liters.
2. It should operate with saturated steam under pressure of 15 to 22 PSI (adjustable).
3. It should have sturdy double wall construction with boiler made of stainless steel at least 16-gauge sheet.
4. Should provide heat resistant SILICON door gasket withstand upto 140 Deg. C.
5. Outer shell should also be made of stainless steel.
6. Stand shall be made of Mild steel with anticorrosion paint.
7. Boiler and outer shell should be provided with air insulation.
8. Lid should be made of steel sheet and tightened by radial locking.
9. Joint-less gaskets should be made of Neoprene rubber.
10. Sterilizer should be fitted with water level arrangement to indicate water position inside the boiler.
11. Pressure gauge, air/steam release cock and two safety valves should be provided.
12. It should have immersion type heating element
13. It should be supplied with cord and plug to work on 220 Volts 50 cycles A.C. supply.
14. Safety features of sterilizers: Door locking facility, Low water protection system, Pressure cut off facility and all other necessary safety features.
15. The unit should be manufactured as per IS specifications Mark ISI:3829 and also should bear the certification.
16. The sterilizer should have minimum 5 years warranty.
Three End User performance certificate must be attached in the technical bids.
AMC and CMC prices must be quoted separately for next 5 years after completion of five year warranty.
Spares/ Consuambles required for the instrument must be quoted and shall be valid for two years.