## Sheet1
Specifications for flow cytometer | Criteria Matching (Yes/No) | Remarks
1. The system must be equipped with 4 solid state lasers (405nm Violet, 488nm Blue & 633-642nm Red, 561nm Yellow-Green) having a minimum 12 fluorescence parameter output to support the multicolor capability of the instrument & field upgradable to 16 colors. All the lasers & detectors should be flexible aligned for data consistency and reliability on a day-to-day basis. The machine should be upgradable with true UV(355 nm) laser in future.
2. Fluorescence detector channel must be configured with PMT for better signal to noise (S/N) ratio and superior resolution of dim intensity rare cell population discovery.
3. All the lasers should be spatially separated & Height, Area, Width information available for all parameters simultaneously to be able to discriminate doublets based on size, granularity & nucleic acid content.
4. Digital signal processing should allow threshold to be set on all available channels simultaneously in any combination of all available parameters during sample acquisition.
5. For high throughput the analysis speed should be at least 35,000 events per second or better with all the parameters available. The system apart from offering low, medium & high flow rates, should also offer high sensitivity fluidics aspiration mode, which can result in higher fluorescence signal resolution for dim stained population.
6. The System should have sample carry over of <0.1% with cells or better for rare cell populations discovery and novel marker identification.
7. The instrument should not need daily compensation to be created by the user. The compensation once done should be valid for reasonable time (preferably 30 days or more) and should be automatically adjusted during daily QC. The system software should be able to do single fluorochrome addition to an existing setting from panel of reagents and recalculate the spillover matrix by running a single tube.
8. The system should not have any event acquisition limit and should enable resolution of <0.2-μm beads from noise in both SSC and/or FSC detector channels from any laser for particle analysis. Sample loading tubes should be open to third party supplier. It would be better if it is not restricted to only 5 ml tubes.
9. This system must be capable of standardization and collaboration between inter-lab/intra-lab through assay portability feature to maintain consistency in data quality.
10. Instrument software must be capable of exporting data as FCS file per population. In addition, system software must allow data overlay of tubes in the same experiment & different experiment. FCS file format should be fcs 3.0 & 3.1. Software should be compatible with FlowJo for data analysis.
11. The system should be in a single tube acquisition format & must be upgradable in future to universal plate and tube loader platform which can accommodate minimum 30 tubes rack as well as 96/384 well plates directly for complete walk away automation.
12. Compatible UPS for at least 30 min standby time of the supplied flow cytometer
13. Twice per year training / workshops for the users to be arranged after equipment installation.
14. Warranty Five year from the date of successful installation in the lab.
15. Sheath fluid, FACS Clean, FACS Shut down sloution etc should be provided (20 Boxes each) along with startup kits and PM Kits (to last for 2 years).
16. A separate Workstation with high RAM (64 GB) with Hard Disk (2TB) and printer should be provided for offline data analysis.
17. Three End User performance certificate must be attached in the technical bids for similar quoted model.
18. AMC and CMC prices must be quoted separately for next 5 years after completion of three year warranty.
Optional Items to be quoted seperately
355 (UV) nm
All consumables apart from the above mentioned required to run flow cytometer should be quoted seperately in USD and quote should be valid for next two years.