## Sheet1
Specifications | Creteria Matching (Yes/No) | Remarks
System should be compatible for algal, microbial, fungal and yeast sample lyophilization.
System should be a compact bench top design with a small footprint and capable of handling low eutectic point samples.
Should have an upright PTFE coated (Teflon coated) stainless steel collector coil condenser with Ice Holding capacity 4 -5 Liters of ice before defrosting.
The system should be capable of removing 2.5 -3 liters within 24 hrs.
It should have dual 1/3 hp or more HCFC/CFC-free refrigeration system to cool collector to >= -105°C or below.
Should have a Moisture sensor to prevent vacuum pump start up when moisture is detected in the collector chamber.
Should have Vacuum control option to maintain set point vacuum level.
Should have hot gas Deforest system.
The system should have a special arrangement for preventing the oil back-streaming into the collector chamber during the power outage. (Vacuum Break Valve)
Should have Touch Screen LCD that displays system operating parameters, set Up parameters and alarm messages for various Parameters like Power failure , Improper line voltage supply, Collector temperature rise above –40° C, Service vacuum pump, Moisture in collector.
The system should have a USB Port to transmit data to a user supplied computer. It should have Auto Emailer option in case of Alarm.
The system should be supplied with a Stainless-steel Drum drying manifold with minimum 12 Ports to connect flasks of different sizes complete with adaptors/ valves.
A Hybrid/ Combination Vacuum Pump with the capacity of 90-100 lit/min designed to use with corrosive chemical like TFA/Acetonitrile /TFA/HNO3 etc. Includes all necessary accessories like oil, power switch, power cord & plug etc.
Tube Holder for 15 ml Tube and 2 ml Tubes with Flask of 750 ml Cylindrical Flask (2 Nos.)
Cylindrical Flask of 40 ml and 150 ml with adaptors having capacity of 8 to 10 falcons in each.
Other important accessories such as flasks, adapters, ampoules, serum bottles, three tier product shelves for tray drying etc.
Power Requirement -230 Volts 50 Hz
Should provide warranty and CMC for 5 Years.