## Sheet1
Date:14/03/2016
Quotations are invited for “Multi-Color Bench-Top FACS Sorter under Buy back of Existing BD Aria Sorter"
Sealed quotations clearly marked “Multi-Color Bench-Top FACS Sorter under Buy back of Existing BD Aria Sorter. ” in two separate bids "Technical and Price offer" along with terms and conditions, warranty should be addressed to:
The Procurement Manager ICGEB Aruna Asaf Ali Marg, New Delhi 110 067
Quotations must reach latest by "20-03- 2017 By 5:00pm
Specification for Multi-color bench-top FACS sorter
S.no | Specification for Multi-color bench-top FACS sorter | Criteria Matching (Please specify YES/NO) | Remarks
1 | Lasers:
2 | 488nm Blue Laser
3 | 633nm Red Laser
4 | 405 nm Violet Laser
5 | 561 nm Yellow Green Laser
6 | Number of parameters: 16 (FSC, SSC and 14 colors)
7 | Configurable up to 18 Fluorescence Detectors and 20 parameters
8 | Fixed align optics with optical arrays of Octagon and trigons
9 | Fluidics cart: includes sheath, waste and cleaning containers
10 | Two and four way sorting into 12x75 mm and 15 mL tubes
11 | Automatic Cell Deposition System (ACDU) for direct sorting on slides, plates of 6, 24, 48, 96 &384
12 | Sorting into microtubes and 12x75-mm tubes
13 | Sample injection chamber for various sample input tubes, including microtubes, 12x75-mm and 15 ml Tubes
14 | Capable to acquire and sort up to 70,000 events/second
15 | Sheath pressure is able to adjust from 5 to 75 psi
16 | Nozzle size 70, 85, 100 and 300 micron sizes.
17 | Integrated with a Bio-safety Cabinet (Class II)
18 | Data Management System, relevant Softwares for acquisition, analysis and sorting
19 | Suitable Workstation (HP4600 or higher, Pent IV at 2.8 GHz or higher), 1 TB Hard-drive, DVD/CD Drive/writer, 19” LED/LCD Monitors (two), External Data Storage drive (Terastation), Laser color printer with two additional cartridges, UPS with 30 minutes back-up
20 | One standalone desktop computer with laser color printer/scanner/copier (27-inch iMac, Retina display).
21 | 2-3 batches/lots of CS&T beads
22 | Accudrop beads
23 | One trained personnel dedicated to the machine
24 | Two free training sessions for students at ICGEB
25 | Sheath fluid for first two years
26 | Warranty:5 years including all spares
27 | Buyback-old BD-Aria sorter
28 | User list with contact details must be attached with Model of the instrument and year of Supply
29 | Certificate must be attached along with the quote by the Vendor stated that Model quoted by the vendor is the latest model , Spares and service will be available for this model for atleast 10 years from the date of installation .
Optional Item to be quoted seperately
1 | Violet laser (405nm) for existing FACS machine