## Sheet1
Specifications | Creteria Matching (Yes/NO) | Remarks
Chest Freezer, -86°C - 2no's
Ultra-Low Temperature
Chest Freezer, -86°C Ultra-Low Temperature, with Foam Insulation, 230V,50Hz
Technical Specification:
Capacity: 550- 660 L
Temp range: Programmable to -86°C in 1°C
Increments, even at ambient temperature up to 32°C
Control: Micro Processor control of temperature and alarms with non-volatile memory
Circulation: Requires fan for reduced power consumption and noise.
Cascade Refrigeration: Hermetically-sealed two stage cascade system with capacity to cope in
High-ambient conditions
Insulation: Polyurethane foam 5"
Lid Latch: Positive, single handed cam action latch with lock
Sub-Lids: Two insulated Fomex panels with stainless steel edging
Security: 4-digit password prevents unauthorized users from changing the temperature set point or
alarm set point.
Power: On-Off switch is located behind the locked panel, preventing power from being accidentally
turned of f.
Battery Backup: Activates alarms and display temperature during power outage.
Alarms: Flexible alarm programming helps avoid false temperature alarms & allows grater user
flexibility in alarm control
Construction:
Interior: stainless steel
Heated Air vent: prevents vacuum formation
Air Filter: Front mounted compressor filter is easily visible and accessible by ¼ turn screw for quick
removal.
Certification: CE and UL certified.
Easy to read LED display with display of system alarms & diagnostics
Dimensions (internal): 30.0" x 57.9" x 23.2"
Warranty: 5 Years
Starting kit to be supplied free of cost for each equipment
Optional Item:
Quote for LN2 Backup system/ CO2 Backup system to maintain temp to -80.c with LN2/CO2