## Sheet1
Autoclave for BSL2 lab
Horizontal Cylindrical Sterilizer -1no
Specifications | Creteria Matching (Yes/No) | Remarks
Horizontal Cylindrical Sterilizer, size: 400mm x 600mm depth, suitable for operation on electricity. Electric load 6KW, capacity 79 ltrs.
The unit will have chamber and door plate of stainless steel 316 quality,
Back plate and steam generator of stainless steel argon-arc welded
with tubular mild steel stand. The jacket would be made of boiler qualitysteel.
Steam working pressure 1.2Kg/cm2 corresponding to working
temperature of 121oC.
The sterilizer would be incorporated with the following mountings
and fittings:
a) Multiport operating valve consisting of 5 positions - off/slow, exhaust, vacuum dry, sterilization and fast exhaust also fitted with one pressure gauge and one compound gauge.
b) Mechanical ejector
c) Self sterilizing vacuum drier
d) One spring loaded safety valve and one vacuum breaker.
e) Chamber discharge line would consist of one plug screen,
one dial thermometer, one swing check valve, one steam trap.
f) Automatic self locking device - The door would have automatic
locking device when it is under pressure.
g) Stainless steel 304 qlty. electric steam generator fitted with 3nos.
2KW heaters each, one low water protection and one automatic
pressure controller, one air break contactor with toggle switch
indicating lamps, gauge glass assembly for the steam generator
would have 2nos. automatic shut off valves. The heaters
would be mounted on a thick stainless steel plate.
h) The jacket would be insulated with glass wool and covered
with stainless steel cover.
The sterilizer will be provided with SPECIAL type of COPPER SENSOR and other related accessories for boiler by which you can save the ELECTRIC and WATER consumption drastically and by which you can save the overall cycle time.
Operating voltage - 440Volt
The unit would be manufactured as per ISI specifications
and also would bear ISI certification Mark IS:3829(Part-I)and
ISO 13485 : 2003 & CE certified.
Warranty: 2 years
Installation of the instrument must include vehicle / labor required for unloading of instrument from Vehicle to site inside campus and to carry old unit under buy back offer from site to vehicle through crane etc.