## Sheet1
Autoclave for BSL3 lab
Horizontal Cylindrical Sterilizer -1no's Under buy back of Old Klenzaid make autoclave 400lts -1no.
Specifications | Creteria Matching (Yes/No) | Remarks
Horizontal Cylindrical Sterilizer, size: 400mm dia x 1100mm depth, Double door, suitable for operation on electricity. Electric load 9KW. Chamber capacity 128 ltrs.
The unit will have chamber and door plate of stainless steel 316 quality, Back plate and steam generator of stainless steel argon-arc welded with tubular mild steel stand. The jacket would be made of boiler quality steel.
Steam working pressure 1.2Kg/cm2 corresponding to working temperature of 121oC.
The sterilizer would be incorporated with the following mountings and fittings:
a) Automatic with pre-selected and variable programs.
b) Mechanical ejector
c) Self sterilizing vacuum drier
d) One spring loaded safety valve and one vacuum breaker.
e) Chamber discharge line would consist of one plug screen, one
swing check valve, one steam trap.
f) Automatic self locking device - The door would have automatic
locking device when it is under pressure.
g) Stainless steel 304 qlty. electric steam generator fitted with 3nos.
3KW heaters each, one low water protection and one automatic
pressure controller, one air break contactor with toggle switch
indicating lamps, gauge glass assembly for the steam generator
would have 2nos. automatic shut off valves. The heaters would be
mounted on a thick stainless steel plate.
h) The jacket would be insulated with glass wool and covered with
stainless steel cover.
The sterilizer will be provided with SPECIAL type of COPPER SENSOR and other related accessories for boiler by which you can save the ELECTRIC and WATER consumption drastically and by which you can save the overall cycle time.
The Heart of the system is PLC based microprocessor with the facility of HMI (Human-Machine-Interface) which is incorporated with the sterilizer.
The digital display at front panel will show the following parameters:
1) Chamber temperature
2) Cycle no.
3) Batch no.
4) Time & Date
5) Alarm indicator
6) Error display
7) Low water indicator
The unit would be manufactured as per Bureau of Indian Standards specifications and also would bear certification mark IS 3829 (Part-I)
ISO 13485 : 2003, CE certified.
Operating voltage - 440Volt
Warranty: 2 years
Installation of the instrument must include vehicle / labor required for unloading of instrument from Vehicle to site inside campus and to carry old unit under buy back offer from site to vehicle through crane etc.
Note:- Kindly visit the site and inspect the location of the autoclave before quoting the tender.