## Sheet1
High pressure Rectangular Autoclave under buy back of old Nat steel Autoclave 24SR/E Capacity 432lts –QTY-2NO’S
S.No | Specification for High Pressure Rectangular Autoclave | Creteria Matching (Yes/No) | Remarks
1 | Horizontal Rectangular High Pressure Sterilizer, with chamber size: 600 x 600 x 1200mm (2’ x 2’ x 4’), single hinged door, suitable for operation on electricity. Electric load 18KW , Capacity 432 ltrs.
2 | The chamber would be hydraulically tested to 2 times the working pressure. The normal working pressure would be 1.2 Kg/cm2 corresponding to temperature of 121oC.
3 | The chamber and door plate would be made of stainless steel AISI 316 quality and electric steam generator would be made of stainless steel AISI 304 quality. The jacket would be made of boiler quality steel.
4 | The unit would be complete with following mountings and fittings:
5 | Automatic with pre-selected and variable programs
a) Mechanical ejectors
b) Self sterilizing vacuum drier.
c) Safety valve spring loaded and vacuum breaker.
d) Pressure and compound gauge
e) Screen plug for chamber discharge line.
f) Chamber discharge line with steam trap and swing check valve.
7 | A stainless steel electric boiler would be provided under the shell with interconnecting steam pipes. The boiler would be complete with 3nos. 6KW immersion heaters mounted on stainless steel plate and electric controls would comprise of an air brake contactor, pressure controller, low water protection.
8 | The boiler would also incorporate a gauge glass assembly and glass tube with protector, water drain and water inlet valves would be provided.
9 | The sterilizer will be provided with SPECIAL type of COPPER SENSOR and other related accessories for boiler by which you can save the ELECTRIC and WATER consumption drastically and by which you can save the overall cycle time.
10 | The Heart of the system is PLC based microprocessor with the facility of HMI (Human-Machine-Interface) which is incorporated with the sterilizer.
11 | The digital display at front panel will show the following parameters:
1) Chamber temperature
2) Cycle no.
3) Batch no.
4) Time & Date
5) Alarm indicator
6) Error display
7) Low water indicator
12 | Printer - Printer that will automatically and continuously monitor and record dates, time of day, load, identification no. and operating parameters i.e. temperature and residence time, throughout the length of the autoclave cycles.
13 | The unit would be manufactured as per Bureau of Indian Standards specifications and also would bear certification mark IS 3829 (Part-I) ISO 13485 : 2003, CE certified.
14 | Warranty: 2 years
15 | Installation of the instrument must include vehicle / labor required for unloading of instrument from Vehicle to site inside campus and to carry old unit under buy back offer from site to vehicle through crane etc.